## Slide 1
Combustion Ion Chromatography
Essential Analytical Technique for
Material Investigation

## Slide 2: Analytical technique are essential
Industrial process are controlled – Analytical results are the indicator of controllability
Analytical Techniques
Pressure drop
Corrosion issue
Fouling
Catalyst deactivation
Water / waste water
Process upset

## Slide 3: Samples and Challenges
Sample can range from
Matrix
Analyte – Trace level
Limited sample size ( 10mg)
Mixture of Organic / inorganic
Gases
C2 to C4
Liquid
Cooling water
DM Water
Solid
mixture

## Slide 4: Ion chromatography
Typical Ion Chromatographic System - Anion Analysis
Eluent
Pump
Guard Column
Analytical Column
Suppressor
Regen In
Conductivity Cell
Chromatography Software
Ion Exchange Separation
Suppressed Conductivity
Data Handling and Instrument Control
Sample Injection
Regent Out
NaOH/ MSA

## Slide 5: Charge Detector (QD)
Detector respond to ionic species at a fixed voltage generating current
Response is proposal to charge – Hence it can be used for universal calibration
Current drawn is directly propositional to ionic strength
Weakly dissociate ions have better response
Compliments the conductivity detector, with additional information of interference & better
sensitivity for weak dissociated ions ( formate, acetate, bromate, pyruvate)
S/N ratio is higher compared to that of conductivity
Dissociation of water is responsible for the background noise
Applied voltage impact the Signal and Noise

## Slide 6: Instrumentation:Combustion IC

## Slide 7: Standard or Sample preparation
Calibration standard
  EDC / Br-cresol / Benzothioazo
  Solvent used –Methanol
  Injected volume -90µL

## Slide 8: Corrosion material
First run
Sample preparation
With SiO2- solid mix

## Slide 9: Water pump deposit
White deposit in the pump was provided to know the type
Sample around 0.2g provided for analysis – insoluble in water
unknown

## Slide 10: Heat exchanger sample
Where the cooling process water side having
Major deposit
Deposit the FeCl3

## Slide 11
SAMPLES with High Conductivity
Anion standard -All seven anion tested
on SPE adsorbed all anions- 5mL
Second Trial on same SPE shows 5 anions
removal of Chloride and Bromide ( 5mL)
Third Trial shows 3 anions > 80% recovery & 50% for Nitrite & Phosphate ( 5mL ) removal of Chloride and Bromide-
Fourth Trial shows 3 anions shows> 90% recovery & 50% for Nitrite & Phosphate removal of Chloride and Bromide
Sample + Anion standard eluted through
SPE
Production sample are expected
Overlay of standards for all four repeat
Standard –Neat
Standard eluted
through SPE
F: 0.5ppm
Cl: 2.5ppm
Br:2.5ppm
NO2-:3.5ppm
NO3: 3.5ppm
SO4: 3.5ppm
PO4: 3.5ppm
Neat standard +
Overlay all repeats
Standard –Neat
Standard -SPE

## Slide 12: Anions in High Ammonia sample
Water
Ammonia
Ammonia
HCl
Air /N2
Ammonia as Matrix
Sample Range: 0.1 - 2%
Requirement is for analysis of trace anion and Cation:
Physical properties :NH3 boils at −33.34 °C (−28.012 °F) at a pressure of one atmosphere, so the liquid must be stored under pressure or at low temperature. Ammonia or ammonium hydroxide is a solution of NH3 in water.
Sample preparation
NH3 (Aq) NH3 (gas)
Vent
Base

## Slide 13
Thank You